# Analytical Method Validation Report

Method Validation Report for SEC-HPLC Purity Method. This report describes the validation of the SEC-HPLC method for determination of monomer purity and size variants (HMW, LMW) of mAb-Val-01 per ICH Q2(R2) guidelines.

Analytical Method: SEC-HPLC for Size Variant Analysis

Method validation of SEC-HPLC method. Method name: SEC-HPLC Purity

# 1. Specificity

Specificity was demonstrated by analyzing placebo, reference standard, and stressed samples. No interference from the placebo matrix was observed.

| Parameter | Result | Criteria |
| --- | --- | --- |
| Placebo interference | None detected | No interference |
| Peak resolution (monomer/HMW) | 2.8 | >= 1.5 |
| Selectivity | Demonstrated | Demonstrated |

# 2. Linearity

Linearity was assessed over the range of 0.1 to 2.0 mg/mL (50-200% of nominal concentration). R2 = 0.9998. Correlation coefficient: 0.9999.

| Parameter | Result | Criteria |
| --- | --- | --- |
| R2 (correlation coefficient) | 0.9998 | >= 0.999 |
| Slope | 1.002 | 0.95-1.05 |
| Y-intercept | 0.15 | < 2.0% |

# 3. Range

The analytical range was confirmed as 0.1 to 2.0 mg/mL, corresponding to 50-200% of the nominal working concentration.

# 4. Accuracy

Accuracy was determined by spiked recovery experiments at three concentration levels (80%, 100%, 120%). Mean recovery: 100.2%. Recovery: 100.2%.

| Parameter | Result | Criteria |
| --- | --- | --- |
| Recovery at 80% | 99.5% | 98-102% |
| Recovery at 100% | 100.2% | 98-102% |
| Recovery at 120% | 100.8% | 98-102% |

# 5. Precision -- Repeatability

Repeatability (intra-day precision) was assessed by 6 replicate injections of the reference standard. %RSD = 0.8%.

| Parameter | Result | Criteria |
| --- | --- | --- |
| RSD (n=6) | 0.8% | <= 2.0% |
| Precision (intra-assay) | 0.8% | <= 2.0% |

# 6. Precision -- Intermediate Precision

Intermediate precision was evaluated across two analysts and three days. Inter-day %RSD = 1.2%.

| Parameter | Result | Criteria |
| --- | --- | --- |
| Inter-day RSD | 1.2% | <= 3.0% |
| Inter-analyst RSD | 1.5% | <= 3.0% |

# 7. LOD and LOQ

The LOD (limit of detection) was determined to be 0.02 mg/mL. The LOQ (limit of quantitation) was determined to be 0.05 mg/mL. These are adequate for detection and quantitation of impurity species.

| Parameter | Result | Unit |
| --- | --- | --- |
| LOD | 0.02 | mg/mL |
| LOQ | 0.05 | mg/mL |

# 8. Robustness

Robustness was evaluated by deliberate variation of column temperature (+/- 2 degC), flow rate (+/- 10%), and mobile phase composition (+/- 2%). The method was found to be robust under all tested conditions.

# 9. System Suitability

System suitability criteria were met in all validation runs. Theoretical plates > 5000, tailing factor < 1.5.

# 10. Summary

The SEC-HPLC purity method has been validated per ICH Q2(R2). All validation parameters meet acceptance criteria. The method is suitable for its intended use in release and stability testing.